## How-To
GAGE R&R / MSA STUDY TEMPLATE
WHAT THIS WORKBOOK DOES
Runs a crossed Gage R&R study with the AIAG Average & Range method for 2–10 parts, 2–3 operators and 2–3 trials. It splits total measurement variation into repeatability (equipment, EV) and reproducibility (appraiser, AV), compares gage variation with part variation and with the tolerance, and returns an ACCEPT / REVIEW / REJECT verdict.
QUICK START
Step | Go to | What to do
1 · Set up the study | Study Setup | Fill the blue/yellow input cells: gage ID, characteristic, tolerance, and the part / operator / trial counts (dropdowns).
2 · Collect readings | Data Collection | Each operator measures every part in random order for every trial, blind to other readings. Enter results in the yellow cells; cells outside your design grey out.
3 · Review results | GRR Calculations | EV, AV, GRR, PV, TV, %Study Variation, %Contribution, %Tolerance, ndc and the overall verdict calculate automatically.
4 · Read the summary | Dashboard | One-page scorecard with verdict, key metrics and three charts. Set up for one-page landscape printing.
5 · Tune the criteria (optional) | Constants Reference | Edit the tolerance multiplier (5.15 or 6) and the acceptance thresholds in the blue cells.
COLOUR CODE
Input |  | Blue text on yellow — type here.
Formula |  | Black text — calculated; do not edit.
Link |  | Green text — pulled from another sheet.
INTERPRETING THE RESULTS
Metric | Criterion | Meaning
%Study Variation | <10% · 10%–30% · >30% | Gage R&R as a share of total observed variation (GRR ÷ TV). Acceptable · marginal · unacceptable bands.
%Tolerance | <10% · 10%–30% · >30% | Gage R&R spread (multiplier × GRR) as a share of the tolerance. Same bands. Can disagree with %Study Variation — check both.
%Contribution | Information only | Variance-based share (GRR² ÷ TV²). GRR and PV contributions add to 100%.
ndc | ≥ 5 | Number of distinct categories the gage can reliably resolve = INT(1.41 × PV ÷ GRR).
EV / AV / GRR | Information only | Equipment (repeatability), appraiser (reproducibility), and combined gage variation, in measurement units.
PV / TV | Information only | Part-to-part variation and total variation (√(GRR² + PV²)) in measurement units.
METHOD NOTES
• Rp is the range of the part averages, where each part average is taken across all operators and trials (AIAG definition).
• Results appear only when every required reading is entered; otherwise the verdict shows INCOMPLETE.
• If the tolerance is blank, %Tolerance is shown as Not assessed and the verdict uses %Study Variation and ndc only.
• Dashboard charts use fixed 10-part ranges, so studies with fewer parts leave empty space on the right of the x-axis.
REFERENCE
Full method and worked example: sixsigmakaizen.com/guides/measurement-system-analysis/
Companion calculator: sixsigmakaizen.com/tools/gauge-rr-calculator/

## Dashboard
GAGE R&R / MSA DASHBOARD
Crossed study · AIAG Average & Range method · 5 parts × 2 operators × 3 trials
STUDY DETAILS |  |  | OVERALL GAGE VERDICT
Gage / instrument | Digital bore gage #14 |  | REJECT
Characteristic | Bore diameter (mm)
Tolerance (USL − LSL) | 0.06 mm |  | Measurement system is not acceptable for this application. Improve the gage or method, then repeat the study.
Study design | 5 parts × 2 operators × 3 trials
Prepared by | —
Study date | —
Data status | Complete (30 of 30 readings)
KEY RESULTS
Metric | Result | Status
%Study Variation (GRR) | 49.47% | Unacceptable
%Tolerance (GRR) | 57.59% | Unacceptable
%Contribution (GRR) | 24.48% | —
Distinct categories (ndc) | 2 | Insufficient
VARIATION COMPONENTS
Component | Value | %Study Var
Repeatability (EV) | 0.00671 | 49.47%
Reproducibility (AV) | 0 | 0%
Gage R&R (GRR) | 0.00671 | 49.47%
Part variation (PV) | 0.011787 | 86.9%
Total variation (TV) | 0.013563 | 100%
READING THE RESULT
Dominant source of gage error: Repeatability (EV) — the gage / equipment itself
Study variation: <10% acceptable, up to 30% marginal. Tolerance: <10% acceptable, up to 30% marginal. ndc ≥ 5.
The two percentage measures can disagree — review both before accepting or rejecting a gage.

## Study Setup
STUDY SETUP
Complete the blue input cells. Part, operator and trial counts resize 'Data Collection', 'GRR Calculations' and the Dashboard automatically.
STUDY IDENTIFICATION
Gage / instrument ID | Digital bore gage #14 | Asset or calibration ID of the gage under study.
Characteristic measured | Bore diameter | The feature being measured.
Unit of measure | mm | Display only (e.g. mm, in, µm).
Prepared by |  | Name of the person running the study.
Study date |  | Date the readings were taken.
SPECIFICATION & STUDY DESIGN
Tolerance (USL − LSL) | 0.06 | Total spec width in measurement units. Required for %Tolerance.
Number of parts (2–10) | 5 | Select from the dropdown. AIAG recommends 10 parts.
Number of operators (2–3) | 2 | Select from the dropdown.
Number of trials (2–3) | 3 | Select from the dropdown.
OPERATORS
Operator 1 name | Operator A | Label used on all sheets and charts.
Operator 2 name | Operator B | Label used on all sheets and charts.
Operator 3 name (only if 3 operators) |  | Ignored unless operators = 3. Defaults to “Operator C”.
STUDY CHECKS
Total readings required | 30
Readings entered | 30
Tolerance specified | OK
Design within AIAG limits | OK
Data entry status | OK — all readings entered
Legend: blue text on yellow = input · black = formula · green = link from another sheet

## Data Collection
DATA COLLECTION
Complete — all 30 readings entered.
Operator / Trial | Part 1 | Part 2 | Part 3 | Part 4 | Part 5 | Part 6 | Part 7 | Part 8 | Part 9 | Part 10
Operator A
Trial 1 | Part 1: 25.001; Part 2: 24.998; Part 3: 25.003; Part 4: 24.999; Part 5: 25.002
Trial 2 | Part 1: 25.002; Part 2: 24.997; Part 3: 25.004; Part 4: 24.998; Part 5: 25.001
Trial 3 | Part 1: 25; Part 2: 24.999; Part 3: 25.002; Part 4: 25; Part 5: 25.003
Average (by part) | Part 1: 25.001; Part 2: 24.998; Part 3: 25.003; Part 4: 24.999; Part 5: 25.002
Range (by part) | Part 1: 0.002; Part 2: 0.002; Part 3: 0.002; Part 4: 0.002; Part 5: 0.002
Operator B
Trial 1 | Part 1: 25.002; Part 2: 24.996; Part 3: 25.005; Part 4: 24.997; Part 5: 25.003
Trial 2 | Part 1: 25; Part 2: 24.999; Part 3: 25.003; Part 4: 24.999; Part 5: 25.001
Trial 3 | Part 1: 25.003; Part 2: 24.998; Part 3: 25.004; Part 4: 24.998; Part 5: 25.002
Average (by part) | Part 1: 25.0017; Part 2: 24.9977; Part 3: 25.004; Part 4: 24.998; Part 5: 25.002
Range (by part) | Part 1: 0.003; Part 2: 0.003; Part 3: 0.002; Part 4: 0.002; Part 5: 0.002
Operator C
Trial 1
Trial 2
Trial 3
Average (by part)
Range (by part)
Part average (all operators) | Part 1: 25.0013; Part 2: 24.9978; Part 3: 25.0035; Part 4: 24.9985; Part 5: 25.002
SUMMARY STATISTICS (feed GRR Calculations)
R-bar — Operator A | Part 1: 0.002; Part 2: Average of the part ranges (repeatability) for this operator.
R-bar — Operator B | Part 1: 0.0024
R-bar — Operator C | Part 2: Used only when 3 operators.
R-double-bar (average of operator R-bars) | Part 1: 0.0022; Part 2: Drives EV (equipment variation).
X-bar — Operator A | Part 1: 25.0006; Part 2: Grand average of this operator's readings.
X-bar — Operator B | Part 1: 25.0007
X-bar — Operator C | Part 2: Used only when 3 operators.
X-diff (range of operator averages) | Part 1: 6.6667e-05; Part 2: Drives AV (appraiser variation).
Rp (range of part averages, all operators) | Part 1: 0.0056667; Part 2: Range of per-part averages pooled across operators (AIAG definition). Drives PV.
DATA COMPLETENESS
Readings required (parts × operators × trials) | Part 1: 30
Readings entered (within study design) | Part 1: 30
Data complete? | Part 1: TRUE

## GRR Calculations
GAGE R&R CALCULATIONS
AIAG Average & Range method. Every cell on this sheet is a formula — inputs live on 'Study Setup' and 'Data Collection'.
STUDY PARAMETERS
Trials | 3
Operators | 2
Parts | 5
Tolerance (USL − LSL) | 0.06
Data complete? | Yes
AIAG CONSTANTS
K1 (by trials) | 3.05 | Lookup on Constants Reference
K2 (by operators) | 3.65 | Lookup on Constants Reference
K3 (by parts) | 2.08 | Lookup on Constants Reference
D4 (R-chart UCL factor, by trials) | 2.575 | Lookup on Constants Reference
INPUTS FROM DATA COLLECTION
R-double-bar (average range) | 0.0022 | Average of operator R-bars
X-diff (range of operator averages) | 6.6667e-05 | Max − min of operator grand averages
Rp (range of part averages) | 0.0056667 | Max − min of part averages (all operators)
R-chart UCL (R-double-bar × D4) | 0.005665 | Upper control limit for the range chart
GAGE R&R RESULTS
Component | Value | %Study Variation | %Contribution | %Tolerance
EV — Repeatability (R-double-bar × K1) | 0.00671 | 49.47% | 24.48% | 57.59%
AV — Reproducibility | 0 | 0% | 0% | 0%
GRR — Gage R&R = √(EV² + AV²) | 0.00671 | 49.47% | 24.48% | 57.59%
PV — Part variation (Rp × K3) | 0.011787 | 86.9% | 75.52% | 101.2%
TV — Total variation = √(GRR² + PV²) | 0.013563 | 100% | 100% | 116.4%
ASSESSMENT
Metric | Result | Verdict | Criterion
%Study Variation (GRR ÷ TV) | 49.47% | Unacceptable | < 10% acceptable · 10%–30% marginal · > 30% unacceptable
%Tolerance (GRR × multiplier ÷ tolerance) | 57.59% | Unacceptable | < 10% acceptable · 10%–30% marginal · > 30% unacceptable
ndc (distinct categories) | 2 | Insufficient | ≥ 5 required
OVERALL GAGE VERDICT | REJECT
Dominant source of gage error | Repeatability (EV) — the gage / equipment itself
Note: %Study Variation and %Tolerance can disagree. A gage that is tight relative to tolerance but loose relative to the observed part variation (or vice versa) will show different verdicts — review both before accepting or rejecting a gage.

## Chart Data
CHART DATA
Auto-linked feed for the Dashboard charts — do not edit. #N/A is intentional: it stops charts plotting unused parts, operators or incomplete data.
1 · COMPONENTS OF VARIATION
Component | %Contribution | %Study Variation | %Tolerance
Gage R&R | 24.48% | 49.47% | 57.59%
Repeatability | 24.48% | 49.47% | 57.59%
Reproducibility | 0% | 0% | 0%
Part-to-Part | 75.52% | 86.9% | 101.2%
2 · RANGE CHART BY PART
Part | Operator A | Operator B | Operator C (n/a) | R-bar (avg range) | UCL
Part 1 | 0.002 | 0.003 | #N/A | 0.0022 | 0.005665
Part 2 | 0.002 | 0.003 | #N/A | 0.0022 | 0.005665
Part 3 | 0.002 | 0.002 | #N/A | 0.0022 | 0.005665
Part 4 | 0.002 | 0.002 | #N/A | 0.0022 | 0.005665
Part 5 | 0.002 | 0.002 | #N/A | 0.0022 | 0.005665
| #N/A | #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A
3 · PART AVERAGES BY OPERATOR
Part | Operator A | Operator B | Operator C (n/a) | Part avg (all ops)
Part 1 | 25.001 | 25.0017 | #N/A | 25.0013
Part 2 | 24.998 | 24.9977 | #N/A | 24.9978
Part 3 | 25.003 | 25.004 | #N/A | 25.0035
Part 4 | 24.999 | 24.998 | #N/A | 24.9985
Part 5 | 25.002 | 25.002 | #N/A | 25.002
| #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A

## Constants Reference
CONSTANTS REFERENCE
AIAG lookup tables and assessment parameters. Blue cells (right) are editable assessment parameters; the K and D4 tables are fixed method constants.
K1 — by number of trials |  | Assessment parameters
Trials | K1 | Parameter | Value | Note
2 | 4.56 | Tolerance multiplier (σ spread) | 5.15 | 5.15 = 99% spread (AIAG MSA 3rd ed.). Use 6 for AIAG MSA 4th ed. / 99.73%.
3 | 3.05 | ndc factor | 1.41 | √2; ndc = INT(1.41 × PV / GRR).
|  | %Study Variation — acceptable below | 10% | Below this: Acceptable.
K2 — by number of operators |  | %Study Variation — marginal up to | 30% | Between the two limits: Marginal. Above: Unacceptable.
Operators | K2 | %Tolerance — acceptable below | 10% | Below this: Acceptable.
2 | 3.65 | %Tolerance — marginal up to | 30% | Between the two limits: Marginal. Above: Unacceptable.
3 | 2.7 | Minimum ndc | 5 | ndc at or above this is considered adequate resolution.
K3 — by number of parts
Parts | K3
2 | 3.65
3 | 2.7
4 | 2.3
5 | 2.08
6 | 1.93
7 | 1.82
8 | 1.74
9 | 1.67
10 | 1.62
D4 — R-chart UCL factor, by trials
Trials | D4
2 | 3.267
3 | 2.575
Source: standard AIAG MSA (Average & Range method) K1/K2/K3 constants.
Source: standard Shewhart control-chart D4 constants; R-chart UCL = R-double-bar × D4.